| Tag No. «Name» - «Description» |  |  |  |  |
| --- | --- | --- | --- | --- |
| Subsystem: «ProcessBreakdownIdentifier» - «ProcessBreakdownName» |  |  |  |  |
| Manufacturer. | «ManufacturerName» | Model No. |  | «OEMPartNumber» |
| Field to record Serial No. | «SerialNumber» |  |  |  |
| Drawing No. | «DocumentList» | Task No. | «TaskName» |  |
| Additional Information: | «Comments» |  |  |  |

| Item | Description | OK | N/A | Punch |
| --- | --- | --- | --- | --- |
|  | Confirm all components have been inspected and are free from damage. |  |  |  |
|  | Confirm antenna faces are free from distortion or cracks. |  |  |  |
|  | Confirm analyser cabinet is free from any transport damage. |  |  |  |
|  | Confirm analyser installation is as per MA-0000-M-1068 General Arrangement drawing. |  |  |  |
|  | Confirm analyser installed at correct distance from belt weigher: Record distance: m. |  |  |  |
|  | Confirm antennas positioned on conveyor belt centreline. |  |  |  |
|  | Confirm antennas positioned in middle of idler bay. |  |  |  |
|  | Confirm antenna faces are parallel to the plane in which the conveyor runs. |  |  |  |
|  | Confirm antennas bolted in position, not welded & mounting is sturdy & rigid. |  |  |  |
|  | Confirm transmit antenna fitted at the correct height: Record distance: mm. |  |  |  |
|  | Confirm receive antenna is fitted at the correct distance below the underside of belt when conveyor is under normal load: Record distance: mm. |  |  |  |
|  | Confirm no obstructions between antennas and metallic reflection points near antennas kept to minimum. |  |  |  |
|  | Confirm analyser cabinet positioned as per vendor specifications. |  |  |  |
|  | Confirm cabinet mounted using bolts, not welded. |  |  |  |
|  | Confirm all bolted connections are of the correct length, grade and tensioned correctly. |  |  |  |
|  | Confirm guarding is installed and secure. |  |  |  |
|  | Confirm all identification and signage is installed as per design drawings and project specifications. |  |  |  |

| COMMENTS, OBSERVATIONS & ATTACHMENTS (Complete a Punchlist Form and attach if any of the checks above marked as “Punch”) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| COMPLETED |  |  |  |  |  |
| I hereby confirm the work has been completed in accordance with Contract specified drawings, specifications & standards. |  |  |  |  |  |
| PERFORMED BY CONTRACTOR: |  | REVIEW BY SUPERVISOR/LEAD ENGINEER: |  | EPCM ACCEPTED BY: |  |
| Signature: |  | Signature: |  | Signature: |  |
| Name: |  | Name: |  | Name: |  |
| Title: |  | Title: |  | Title: |  |
| Date: |  | Date: |  | Date: |  |